## Run 1
Lab 104D SierraTherm CVD Coater Data Sheet
Angelica & Doug
Operator(s): | [6 empty] | Sample ID: | 20250708 |  |  |  | Run#: | 1
Date: | 2025-07-08 | [5 empty] | Glass Type: | MSVD Starphire 2.5 mm
Pre-Coating Check List | [10 empty] | Conveyor Furnace Setpoints
standard.rcp(Standard)
Furnace Setpoints | X | Cart Air ON | X | Vaporizer Setpoints | X | Sierra Therm Exhaust ON |  |  | X |  | Zone 1 | 520
Carrier Air Setpoints | X | Belt Speed Correct | X | Coater Inlet Setpoint | X | Condenser Temperature |  |  | X |  | Zone 2 | 570
Carrier/CP N2 Setpoints | X | Chemistry Setpoints | X | Bypass Setpoint | Off | Coater Temperature |  |  | X |  | Zone 3 | 620
Blower ON | X | Lines Full | X | Coater Purge Setpoint | X | [5 empty] | Zone 4 | 640
Chilled Water ON | X | Syringes Set to Coat | X | Temp. Controllers On | X | [5 empty] | Zone 5 | 660
Control Panel On | X | Coater Cooling ON | X | Coater Exhaust Setpoint | X | [5 empty] | Zone 6 | 670
[11 empty] | Zone 7 | 680
[11 empty] | Zone 8 | 690
Conveyor Furnace Settings | [10 empty] | Zone 9 | 660
Recipe: |  |  |  |  | N2 Flows | [5 empty] | Zone 10 | 660
cvd-6351stslot.rcp | X |  |  | Entry Curtain |  | 200 |  |  |  |  | Zone 11 | 620
doe.low2a.rcp |  |  |  | Burnout Section |  | 100
doe.low2a-20C.rcp |  |  |  | Purge N2 |  | 50
Belt Speed (Inches/Min.) |  |  |  | Firing Section |  | 120
Run In | 5 |  |  | Exit curtain |  | 200
Coating | 5 |  |  | Furnace Exhaust Pressure |  | 0.03
Run Out | 5
Cart A | [17 empty] | Cart B
|  | Initial | 1 | 2 | 3 | [5 empty] | Condenser Temperature © |  |  | Coater Temperature (C) | [5 empty] | Initial | 1 | 2 | 3 | [5 empty] | Condenser Temperature (C) |  | Coater Temperature (C)
Exhaust Blower (Hz) |  | 6 | 6 | 5.9 | 6 | [5 empty] | Inlet | 130.4 |  | TL |  |  |  | Exhaust Blower (Hz) | [10 empty] | Inlet |  | TL
Exhaust Flow (Δ): |  | 46.2 | 46.9 | 46 | 45.9 | [5 empty] | Outlet | 29.9 |  | ML |  |  |  | Exhaust Flow (Δ): | [10 empty] | Outlet |  | ML
[14 empty] | BL | [16 empty] | BL
| Coater Inlet Mixer: | Setpoint | Actual |  | Coater Inlet Line: | Setpoint | Actual |  | Bypass Temp: | Setpoint | Actual |  |  | TR |  |  |  |  | Coater Inlet Mixer: | Setpoint | Actual |  | Coater Inlet Line: | Setpoint | Actual |  | Bypass Temp: | Setpoint | Actual |  | TR
|  | 160 | 160 |  |  | 160 | 156 |  |  | 150 | 150 |  |  | MR | [5 empty] | 160 |  |  |  | 160 |  |  |  | 150 |  |  | MR
[14 empty] | BR | [16 empty] | BR
Chemistry | [17 empty] | Chemistry
Cell # | Precursor | Flow (ml/hr) | MFC max | Carrier Gas | MFC% | Gas Flow (SLM) |  | Vaporizer Target Temp (C) | N2 Heater Temp. (C) | Vaporizer Actual Temp. (C) | Valve Temp. (C) | Line Temp. (C) |  | TFE |  |  |  | Cell # | Precursor | Flow (ml/hr) | MFC max | Carrier Gas | MFC% | Gas Flow (SLM) |  | Vaporizer Target Temp (C) | N2 Heater Temp. (C) | Vaporizer Actual Temp. (C) | Valve Temp. (C) | Line Temp. (C)
V1 | None |  | 50SLM | N2 |  | N/A |  | 177 | ----- | ----- | ----- | ----- |  |  | TFE Oil Jacket (C) | TFE Line 1 (C) | TFE Line 2 (Flexline) (C) | V1 | None |  | 50SLM | N2 |  | N/A |  | 177 | ----- | ----- | ----- | -----
V2 | TFA | 3.5 | 50SLM | N2 |  | 14 |  | 177 | 177 | 180 | 150 | 177 |  | Setpoint | ---- | ---- | ---- | V2 | TFA |  | 50SLM | N2 |  |  |  | 177
V3 | Mix | 23 | 30SLM | N2 | ----- | 14 |  | 177 | 165 | 164 | 150 | 181 |  | Actual |  |  |  | V3 | MBTC |  | 30SLM | N2 | ----- |  |  | 177
V4 | Water | 3.5 | 30SLM | N2/O2 Mix |  | 30 |  | 177 | 120 | 120 | 150 | 177 | [5 empty] | V4 | Water |  | 30SLM | N2/O2 Mix |  |  |  | 177
CP | None | -------- | 200SLM | N2 |  | 58 |  | 177 | 177 | -------- | 146 | 150 | [5 empty] | CP | None | -------- | 200SLM | N2 |  |  |  | 177 |  | --------
[5 empty] | Total Vaporizer Carrier Gas Flow: | [17 empty] | Total Vaporizer Carrier Gas Flow:
Initial Comments: Vap 4: 24 N2, 6 O2 FTO on MSVD Coat Time: Header/Trailer on Refractory, sample on boro |  |  |  |  | Final Comments: | [12 empty] | Initial Comments: |  |  |  |  | Final Comments:
Vap 4: N2 24 L/min, O2 6 L/min Mix: 95% MBTC:5% DMTC Sample, header, and trailer on boro |  |  |  |  | The MBTC line was flooded by adding 40 ml instead of 10 ml when filling the line

## Run 2
Lab 104D SierraTherm CVD Coater Data Sheet
Angelica & Doug
Operator(s): | [6 empty] | Sample ID: | 20250708 |  |  |  | Run#: | 2
Date: | 2025-07-08 | [5 empty] | Glass Type: | MSVD Starphire 2.5 mm
Pre-Coating Check List | [10 empty] | Conveyor Furnace Setpoints
standard.rcp(Standard)
Furnace Setpoints | X | Cart Air ON | X | Vaporizer Setpoints | X | Sierra Therm Exhaust ON |  |  | X |  | Zone 1 | 520
Carrier Air Setpoints | X | Belt Speed Correct | X | Coater Inlet Setpoint | X | Condenser Temperature |  |  | X |  | Zone 2 | 570
Carrier/CP N2 Setpoints | X | Chemistry Setpoints | X | Bypass Setpoint | Off | Coater Temperature |  |  | X |  | Zone 3 | 620
Blower ON | X | Lines Full | X | Coater Purge Setpoint | X | [5 empty] | Zone 4 | 640
Chilled Water ON | X | Syringes Set to Coat | X | Temp. Controllers On | X | [5 empty] | Zone 5 | 660
Control Panel On | X | Coater Cooling ON | X | Coater Exhaust Setpoint | X | [5 empty] | Zone 6 | 670
[11 empty] | Zone 7 | 680
[11 empty] | Zone 8 | 690
Conveyor Furnace Settings | [10 empty] | Zone 9 | 660
Recipe: |  |  |  |  | N2 Flows | [5 empty] | Zone 10 | 660
cvd-6351stslot.rcp | X |  |  | Entry Curtain |  | 200 |  |  |  |  | Zone 11 | 620
doe.low2a.rcp |  |  |  | Burnout Section |  | 100
doe.low2a-20C.rcp |  |  |  | Purge N2 |  | 50
Belt Speed (Inches/Min.) |  |  |  | Firing Section |  | 120
Run In | 5 |  |  | Exit curtain |  | 200
Coating | 5 |  |  | Furnace Exhaust Pressure |  | 0.03
Run Out | 5
Cart A | [17 empty] | Cart B
|  | Initial | 1 | 2 | 3 | [5 empty] | Condenser Temperature (C) |  |  | Coater Temperature (C) | [5 empty] | Initial | 1 | 2 | 3 | [5 empty] | Condenser Temperature (C) |  | Coater Temperature (C)
Exhaust Blower (Hz) |  | 6 | 5.9 | 6 | 6 | [5 empty] | Inlet | 130 |  | TL |  |  |  | Exhaust Blower (Hz) | [10 empty] | Inlet |  | TL
Exhaust Flow (Δ): |  | 45.4 | 45.7 | 46.1 | 46.4 | [5 empty] | Outlet | 30.9 |  | ML |  |  |  | Exhaust Flow (Δ): | [10 empty] | Outlet |  | ML
[14 empty] | BL | [16 empty] | BL
| Coater Inlet Mixer: | Setpoint | Actual |  | Coater Inlet Line: | Setpoint | Actual |  | Bypass Temp: | Setpoint | Actual |  |  | TR |  |  |  |  | Coater Inlet Mixer: | Setpoint | Actual |  | Coater Inlet Line: | Setpoint | Actual |  | Bypass Temp: | Setpoint | Actual |  | TR
|  | 160 | 160 |  |  | 160 | 156 |  |  | 150 | 150 |  |  | MR | [5 empty] | 160 |  |  |  | 160 |  |  |  | 150 |  |  | MR
[14 empty] | BR | [16 empty] | BR
Chemistry | [17 empty] | Chemistry
Cell # | Precursor | Flow (ml/hr) | MFC max | Carrier Gas | MFC% | Gas Flow (SLM) |  | Vaporizer Target Temp (C) | N2 Heater Temp. (C) | Vaporizer Actual Temp. (C) | Valve Temp. (C) | Line Temp. (C) |  | TFE |  |  |  | Cell # | Precursor | Flow (ml/hr) | MFC max | Carrier Gas | MFC% | Gas Flow (SLM) |  | Vaporizer Target Temp (C) | N2 Heater Temp. (C) | Vaporizer Actual Temp. (C) | Valve Temp. (C) | Line Temp. (C)
V1 | None |  | 50SLM | N2 |  | N/A |  | 177 | ----- | ----- | ----- | ----- |  |  | TFE Oil Jacket (C) | TFE Line 1 (C) | TFE Line 2 (Flexline) (C) | V1 | None |  | 50SLM | N2 |  | N/A |  | 177 | ----- | ----- | ----- | -----
V2 | TFA | 3.5 | 50SLM | N2 |  | 14 |  | 177 | 177 | 177 | 150 | 177 |  | Setpoint | ---- | ---- | ---- | V2 | TFA |  | 50SLM | N2 |  |  |  | 177
V3 | Mix | 23 | 30SLM | N2 | ----- | 14 |  | 177 | 165 | 165 | 150 | 177 |  | Actual |  |  |  | V3 | MBTC |  | 30SLM | N2 | ----- |  |  | 177
V4 | Water | 3.5 | 30SLM | N2/O2 Mix |  | 30 |  | 177 | 120 | 120 | 150 | 177 | [5 empty] | V4 | Water |  | 30SLM | N2/O2 Mix |  |  |  | 177
CP | None | -------- | 200SLM | N2 |  | 58 |  | 177 | 177 | -------- | 145 | 150 | [5 empty] | CP | None | -------- | 200SLM | N2 |  |  |  | 177 |  | --------
[5 empty] | Total Vaporizer Carrier Gas Flow: | [17 empty] | Total Vaporizer Carrier Gas Flow:
Initial Comments: Vap 4: 24 N2, 6 O2 FTO on MSVD Coat Time: Header/Trailer on Refractory, sample on boro |  |  |  |  | Final Comments: | [12 empty] | Initial Comments: |  |  |  |  | Final Comments:
Vap 4: N2 24 L/min, O2 6 L/min Mix: 95% MBTC:5% DMTC Sample on boro, Header and trailer on refractory |  |  |  |  | started chemistry ~3 min late due to Billco washer water tank leak into 104D